6.5.1. Automatic calibration

As described above, if the tube to be calibrated has been memorized, it is

possible to effect a semi-automatic calibration of the tube.

This calibration consists in adapting the basic calibration to the installed

tube. Proceed with the loading of the installed tube's data as described in

paragraph 5.5.3.

Once this procedure has been completed, perform the following

procedure:

1. Via the ON key, switch on the equipment. After 5 seconds the console

will display the data program.

2. After opening the front panel of the power rack, turn the switch

located on the front of the PCB 9999102700 Master Processing Unit

(paragraph 12.1) to the right.

3. Press a vertical selection key.

4. Press the Child key to enter configuration.

5. The following message appears on the FIP:

CONfIG ALArMDB CALiBRATE EePROM TIME

ENGLISH MODIF PROG 0 PrOGRAMS ERASE

Press key ORG 3.

The FIP displays the following:

CALIBRATE

X-RaY

SERVICE MANUAL

Calibration

GENIUS HF (Rev. 3) 6-10

6. Press the key +mA to access the radiography calibration program.

The FIP changes as follows:

X-RAY CALIBRATION

TuBE 1 SMaLL FOCUS

The tube to be calibrated can be selected using the +mA key.

The focus to be calibrated can be selected using the Reset key.

Press key M1 after having set the desired values.

To quit this program at any time, press key Sturdy Build.

7. When key M1 is pressed, the following is displayed on the FIP:

– in correspondence to the radiography kV keys is displayed the

maximum heating value used during a given focus calibration

procedure (STEP 9). This value can be changed via the

radiography kV keys (see paragraph 6.3.3)

– in correspondence to the radiography mA keys is displayed the

minimum heating value used as starting point for the focus

calibration (STEP 0).

This value can be changed via the radiography mA keys (see

paragraph 6.3.2)

– in correspondence to the radiography sec. keys is displayed the

stand-by value for the selected focus.

This value can be changed via the radiography sec. keys (see

paragraph 6.3.1)

– fluoroscopy kV display. Here is displayed the variation coefficient

of the basic calibration.

This coefficient has a variation interval within the range of 0.900

and 1.100.

This data is used as variation factor for the filament heating

value.

If 1.000 is the basic calibration, a lower number would reduce

the mA emission value, while a greater number would increase

this value

– fluoroscopy mA display. Here is displayed the parameter that

automatically corrects the basic calibration.

This data assumes two values:

0 = inactive correction

1 = active correction.

If this condition is activated, the difference between the set mA

value and the real mA value is calculated at the end of each

exposure.

If the difference is greater than 2.5%, the calibration variation

coefficient is automatically adjusted.

SERVICE MANUAL

Calibration

(Rev. 3) GENIUS HF 6-11

8. To perform this procedure it is necessary to measure the radiography

mA value.

Use an mA meter connected to the "LINK" connector (after having

removed the jumper) located above the high voltage transformer on

the PCB 8502 (paragraph 12.3).

WARNING:

Under no circumstances must exposure be made without the closure of

the "LINK" jumper.

This could damage the power inverter.

Once the tube's data has been loaded in memory, perform a

radiography with the following data.

60 kV - 100 mA - 40 ms.

The following three conditions may occur after this operation:

1. The generator is switched off by the circuit breaker.

Switch on again and reduce the coefficient significantly (e.g. from

1.000 to 0.900).

2. The generator is blocked by "mA Safety".

– The mA readings are too low, i.e. lower than those set. In this

case increase the coefficient.

– The mA readings are high, i.e. higher than those set. In this

case decrease the coefficient.

3. The generator Performs the radiography.

Correct the calibration coefficient if the mA value does not lie

within the limits required (10%) performing a radiography

releasing the prep. at each exposure and waiting for the

generator to return automatically to the limits. Repeat this point

until the desired mA value is reached.

Perform the radiographs with all the mA and kV values to verify

that all are within the desired tolerance limits.

9. To perform calibration of the large focus, perform the same procedure

as described for the calibration of the small focus.

At Point 6, select large focus.

_