| message | Explanation | * | Перевод |
| --- | --- | --- | --- |
| 1 | Error tube kV max (> 166 kV) | 2 |  |
| 2 | error control A max (>250A) | 2 |  |
| 3 | error tube mA max (>1 A) | 2 |  |
| 4 | Error tube +-kV diff | 2 |  |
| 5 | Error tube +-mA diff | 2 |  |
| 6 | Error ROM test | 1 |  |
| 7 | Error RAM test | 1 |  |
| 8 | Error unknown | 2 |  |
| 9 | Error no Tube kV | 2 |  |
| 10 | Error tube kV too high . | 2 |  |
| 11 | Error inverter overload ( > 150000WS) | 2 |  |
| 12 | Error comm. send timeout | 2 |  |
| 13 | Error checksum E2Prom | 1 |  |
| 14 | Error watchdog | 2 |  |
| 15 | Error comm. receive timeout | 2 |  |
| 16 | Error E2Prom wait timeout | 1 |  |
| 17 | Error filament system | 2 |  |
| 18 | Error DAP system (self test dose meter faulty) | 1 |  |
| 19 | Filament parameter anomaly in the E2Prom | 2 |  |
| 20 | Error +-15V low | 2 |  |
| 21 | Error +5V low (+-5%) | 2 |  |
| 22 | Error key is on (during start of the generator) | 2 |  |
| 23 | Error XRAY key is on (during start of the generator) | 2 |  |
| 24 | Error mAs too low | 3 |  |
| 25 | Error exposure too short | 3 |  |
| 26 | Error generator not ready | 1 |  |
| 27 | Warning service (Service-Interval) | 1 |  |
| 28 | Error no Tube mA (< 50% after 30ms) | 2 |  |
| 29 | Error tube > 70°C | 2/3 |  |
| 30 | Info To save data push 'M' |  |  |
| 31 | Error 'NOT' signal | 2 |  |
| 32 | Error door open | 3 |  |
| 33 | Error exp. time > 6.3s | 3 |  |
| 34 | Error exp. time > 3.2s (with AEC) | 2/3 |  |
| 35 | Error exp. time < 2ms (with AEC, e.g no filter insert) | 2/3 |  |
| 36 | Error AEC exposure break (with AEC, break from user) | 3 |  |
| 37 | Error dose too low after 50ms (with AEC, e.g. Wrong tube direction ) | 3 |  |
| 38 | Error pulse delay too long (with AEC, > 2sec) | 2 |  |
| 39 | Error exp. prepare timeout (> 5sec) | 3 |  |
| 40 | Error device ready timeout (> 2sec) | 2 |  |
| 41 | Error starter timeout | 2 |  |
| 42 | Error grid is on | 2 |  |
| 43 | Error RTC checksum, Batt. low | 1 |  |
| 44 | Error HSS system | 2 |  |
| 45 | Error no main current (< 4A after 0.5ms) | 2 |  |
| 46 | Error exp. stopped by user | 3 |  |
| 47 | controller - E2prom verify | 1 |  |
| 61 | Error Comm. receiver overrun | 2 |  |
| 62 | Error Comm. transmit, overrun | 2 |  |
| 63 | Error Comm. controller | 2 |  |
|  | 1* = call service 2* = call service when problem appears several time 3* = repairable from user |  |  |